Startup Procedure for the Perkin Elmer Spinning Disk Confocal

- Turn on Nitrogen valve.

- Sign in on the user login sheet and write down the current clock time under “Begin.”

- Turn on the Line Conditioner (#1).

- Turn on the Power Strip (#2)

- Turn on the laser module (#3). Note: it takes 5 to 10 minutes for the lasers to warm up.

- Turn on the Intensilight epifluorescence light source (#4). Note! Only turn on this light source if you need to visualize epifluorescence through the eyepieces. Note! Turn off Intensilight for overnight time-lapses.

- Turn on the Nikon Ti microscope (#5).

- Turn on the computer (#6).

- Log into your Windows account.

- Double-click on the Volocity software icon to start the software. If you don't see the right side of the screen populated with all of the microscope/camera controls (i.e., it's just a long white rectangle of nothing), close the software, unplug the camera power supply, plug is back in, and restart the software.

Shutdown Procedure for the Perkin Elmer Spinning Disk Confocal

Check Google Calendar to see if someone is signed up to use the microscope after you.

If someone is booked within one hour after you are finished with the microscope…

- Save your images.

- Clean the oil objective lens(es) with Sparkle and lens paper.  Clean the water lens with lens paper.

- Exit the software and log out of your Windows account.

- Fill out the logbook.  Record the clock time in the “End” space of the login sheet.

If someone is booked >1 hr after you are finished with the microscope…

- Save your images.

- Clean the oil objective lens(es) with Sparkle and lens paper.  Clean the water lens with lens paper.

- Exit the software, log out of your Windows account and shutdown the computer (#6).

- Turn off the Nikon Ti microscope (#5).

- Turn off the Intensilight epifluorescence light source (#4).

- Turn off the laser module (#3). Note! Do not turn off the power strips until all the lights on the front of the laser module go out.

- Turn off the Power Strip (#2) and Line Conditioner (#1).

- Fill out the logbook. Record the current time in the “End” space of the login sheet.

- Put the cover on the microscope.

Partial System Restart in Case of a System Malfunction

If you experience a malfunction with the system, you can try performing a partial system shut down.

- Exit the software, log out of your Windows account, and shutdown the computer (#6).

- Turn off the Nikon Ti microscope (#5).

- Turn off the Line Conditioner  #1. Please note: Do NOT turn off the Intensilight (#4) or the power strip labeled #2!